|  |
| --- |
| **Metadata: Mohawk River at Rexford Bridge**C:\HRECOS\HRECOS_logo.small.TIF**Location:** Rexford Bridge, Mohawk River ([42.851 N, 73.887 W](https://maps.google.com/maps?q=42.851132,+-73.887180&hl=en&sll=42.668084,-73.810681&sspn=0.237296,0.357399&t=h&z=17))**Data collection period:** 08/04/2014 – present**Parameters:** acidity, dissolved oxygen, specific conductance, turbidity, water temperature, and water elevation\* (\*USGS). |
| *Disclaimer: HRECOS is a research project. No warranty—either express or implied—is made for any information presented by this program.* |
| **Contacts**:Brittney Flaten, HRECOS Coordinator and Site ManagerNY State Dept. of Environmental Conservation256 Norrie Point Way, Staatsburg, NY 12580Phone: 845-889-4745Email: brittney.flaten [at] dec.ny.gov |
| **Station details:**The HRECOS water quality station at Rexford Bridge is located on the north side of the Mohawk River/Erie Canal, just downstream of the city of Schenectady. Equipment is mounted on the northerly remnants of the Rexford Aqueduct just east of the Rexford Bridge. Preliminary cross-sectional sampling at this site showed no obvious differences across the channel. Sensors are approximately 2 ft. above the riverbed.All parameters are measured using a YSI EXO2 sonde. The following parameters are reported at this site every 15 minutes: acidity, dissolved oxygen (% saturation and mg/L), specific conductance, turbidity, water temperature, and water elevation\* (see the section titled “Sensor Specifications” for more information). Data is logged to a Campbell Scientific CR1000 datalogger and is transmitted hourly to the HRECOS database via a Raven XTV cellular modem.\*Water elevation is measured by a USGS co-located station using an OTT Compact Bubbler System. Water level is converted from NGVD29 to NAVD88 relative to sea level (USGS gage height + 200 ft – 0.538 ft). Only 15-minute data is retained for HRECOS, but the gage records at 5-minute intervals. Original USGS data can be accessed at the following URL: <http://waterdata.usgs.gov/ny/nwis/uv/?site_no=01355475><http://waterdata.usgs.gov/ny/nwis/uv/?site_no=01355475&PARAmeter_cd=00065,00060>.  |
| **Distribution terms:**HRECOS requests that attribution be given whenever HRECOS material is reproduced and re-disseminated and the HRECOS Coordinator be notified prior to publications including any part of the data. Example citation: “Hudson River Environmental Conditions Observing System. 2012. Accessed April 13th, 2016. <http://www.hrecos.org/>.” |
| **Data Quality Assurance:**Data collection and verification have been performed since the establishment of this station, according to the HRECOS Quality Assurance Project Plan, which is available at [www.hrecos.org](http://www.hrecos.org) \*The level gage is maintained by the U.S. Geological Survey. Data QAQC for this parameter occurs on a different schedule than the sonde parameters (~annually, as opposed to quarterly). Verified data may have been corrected based on field measurements, sensor calibrations, sensor cleanings, and other observations using standard USGS methodology.  Unverified data is [provisional and subject to revision](http://waterdata.usgs.gov/ny/nwis/?provisional).  |
| **QAQC Comment Code definitions:** General Errors[GIM] instrument malfunction [GIT] instrument recording error, recovered telemetry data [GMC] no instrument deployed due to maintenance/calibration [GPF] power failure/low battery [GQR] rejected due to QAQC checks[GSM] see metadata [GIC] no instrument deployed due to ice [GNF] deployment tube clogged/no flow [GOW] out of water eventSensor Errors [SBO] blocked optic [STF] catastrophic temperature sensor failure [SCF] conductivity sensor failure [SDF] depth port frozen [SDP] DO membrane puncture [SDO] DO suspect [SIC] incorrect calibration/contaminated standard [SNV] negative value [SPC] post calibration out of range [SSD] sensor drift [SSM] sensor malfunction [SOW] sensor out of water [SSR] sensor removed (not deployed) [STS] turbidity spike [SWM] wiper malfunction/loss Comments (CAB) algal bloom (CAF) acceptable calibration/accuracy error of sensor (CAP) depth sensor in water, affected by atmospheric pressure (CBF) biofouling (CCU) cause unknown (CDA) DO hypoxia < 28 percent saturation (CDB) disturbed bottom (CDF) data appear to fit conditions (CFK) fish kill (CIP) surface ice present at sample station (CLT) low tide (CMC) in field maintenance/cleaning (CMD) mud in probe guard(CND) new deployment begins (CRE) significant rain event (CSM) see metadata (CTS) turbidity spike (CVT) possible vandalism/tampering (CWD) data collected at wrong depth (CWE) significant weather event |

|  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Parameter | Units | Sensor type | Model | Range | Accuracy | Resolution | Response |
| **Acidity** | Hydrogen ion concentration (pH) | Glass combination electrode | 599702 | 0 – 14 units | ±0.1 pH units within ±10°Cof calibration temperature;±0.2 pH units for entire temp range | 0.01 | T63<3 sec |
| **Specific Conductivity** | microsiemens per cm (mS/cm) | 4-electrode nickel | 599870-01 | 0 – 200 mS/cm | 0-100: ±0.5% of reading or 0.001, w.i.g.; 100-200: ±1% of reading | 0.0001 to 0.01, range-dependent | T63<2 sec |
| **Dissolved oxygen** | Air saturation (%)¾¾¾¾mg/L | Optical, luminescence lifetime ¾¾¾¾Calculated | 599100-01 | 0 – 500%¾¾¾¾0 – 50 mg/L | 0 – 200%: ±1%200 – 500%: ±5%¾¾¾¾0 – 20 mg/L: ±0.1 mg/L or 1% (whichever is greater);20 – 50 mg/L: ±-5% | 0.1%¾¾¾¾0.01 mg/L | T63<5 sec |
| **Turbidity** | Formazin Nephelometric Unit (FNU) | Optical, 90° scatter | 599101-01 | 0 – 4000 FNU | 0-999: 0.3 or±2% of reading, whichever is greater; 1000-4000: ±5% of reading | 0.01 FNU | T63<2 sec |
| **Water temperature** | Celsius (°C) | Thermistor | 599870-01 | -5 to +50°C | -5 to 35°C: ±0.01°C35 to 50°C: ±0.05°C | 0.001°C | T63<1 sec |
| **Water elevation** | Feet/Meters | Nitrogen bubbler | OTT CBS | 0 – 50 ft (15.24 m) | 0 – 15 ft: ±0.01 ft (0.003 m);15 – 35 ft: ±0.065%;35 – 50 ft (0.006 m) | Max. traceable rate of change: 3 ft/minute | Vented to atmosphere (no pressure correction needed) |

**Table 1. YSI EXO2 sensor and OTT compact bubbler specifications.**